## Slide 1: GA5000 FAQ’s

## Slide 2: GA5000 FAQ’s – Set Up
Question | Answer
How can I align my charger correctly? | Ensure the notches on both the plug and the socket are correctly aligned. Damaged pins can be difficult to repair and may result in an expensive replacement.
The anemometer that I use with my GA5000 reads in m/s, is it possible to read in m3/hr? | Yes, this can be changed either on the analyser or through gas analyser manager, to read in m3/hr you will have to enter the pipe diameter, to do this: 1. Plug in your anemometer, 2. Power on the analyser, 3. Ensure that the ID you are using has a pipe diameter assigned to it, 4. Press the menu key, select user settings, key 2 for units of measurement, key 2 again to change flow to be measured in m3/hr, 5. Take a flow reading as normal

## Slide 3: GA5000 FAQ’s – Set Up
Question | Answer
How can I download the stored readings from my analyser? | Stored readings can be downloaded using the Gas Analyser Manager software or using the free basic download software on the Geotech website through analyser registration. Please contact sales@qedenv.co.uk for further information.

## Slide 4: GA5000 FAQ’s – Service & Calibration
Question | Answer
How do I perform a zero calibration on a 5000 series analyser? | To zero calibrate a 5000 Series analyser: 1) Power on the analyser and press next to go on to the main read screen 2) Press the options key 3) Press key 2 – gas check and then press key 1 to confirm this 4) Once you are through to the calibration screen, you will notice a range of mixtures on the left-hand side 5) Using the 2 and 8 keys, select ‘fresh air’ so that it is green and press the enter key 6) Let the pump run for 90 seconds 7) Press the left arrow key for ‘check’ 8) Allow that timer to run down 9) Press the middle arrow key for ‘calibrate’ 10) Once that is complete, the CH4 should read zero and the O2 should read 20.9%

## Slide 5: GA5000 FAQ’s – Service & Calibration
Question | Answer
What does a full factory service and calibration consist of? | The full factory service and calibration consist of thorough analyser assessment, service and repair, recalibration, test and verification. Please contact service@qedenv.co.uk for further details.

## Slide 6: GA5000 FAQ’s – Equipment Information
Question | Answer
What are the definitions of pressure? | Relative Pressure – This is the same as static pressure Differential Pressure – The pressure exerted by a gas when the body on which the pressure is exerted is not in motion i.e. the impact pressure of the gas on a fixed pitot tube or orifice plate Static Pressure – This pressure measurement is obtained from the inlet port relative to the outlet port. I.e. the pressure of the monitoring point relative to atmosphere (Barometric Pressure) Barometric Pressure – This is atmospheric pressure at a given location System Pressure – This is the same as static pressure

## Slide 7: GA5000 FAQ’s – Equipment Information
Question | Answer
What do I get when my firmware changes? | Please contact our technical support team for further information on this are we are continuously improving the features and functions technical@qedenv.co.uk
Where can I find product warranty information? | You can find it HERE
Where can I find the firmware update tool for the GA5000, GEM5000 or BIOGAS 5000? | You can find it HERE

## Slide 8: GA5000 FAQ’s – Trouble-shooting
Problem | Answer
I can’t get my USB connections lead to work | The driver’s installation is in two parts and it can be easy to overlook or cancel the second part. Try re-installing the drivers, and check your ‘system properties’ to ensure the COM port is mapped.
I am using Gas Analyser Manager and my data and IDs have disappeared | Please check that your ‘data folder’ is set to the correct location of your files. This can be found in the ‘tools’ menu then ‘options.
The readings from my analyser are not the values that I am expecting | What values are expected? Ensure there is no blockage? When was the last user calibration performed? Is there calibration gas available on site? Is there access to a comparison analyser? This information will assist our technical support team towards quick diagnosis and resolution to the query. Please contact our technical support or take a look at this useful video

## Slide 9: GA5000 FAQ’s – Trouble-shooting
Problem | Answer
My GA5000 doesn’t last long even after a full charge, why? | We first need to check that the charging process is being completed correctly; the instrument must be off, the charger then plugged in and the unit should be left for at least 3 hours. If at any point during these 3 hours the instrument is switched on, then the charge cycle will be interrupted and will not be fully charged. There are a few things you can do to get the most out of the battery life: - – Turning off the back light when it’s not needed, turning off navigation and GPS when it’s not needed – Turning off user prompts if they are not needed

## Slide 10: GA5000 FAQ’s – Trouble-shooting
Problem | Answer
The screen has frozen on my GA5000 or GEM5000 or Biogas 5000, what can I do? | Press and hold the ‘on’ button for approximately 20 seconds. This will reset the hardware. If this does not resolve the issue please contact our technical support team www.technical@qedenv.co.uk