## Slide 1: Biometane 3000 FAQ’s

## Slide 2: Biomethane 3000 FAQ’s
Question | Answer
My instrument keeps displaying a ‘low flow’ warning, why? | This is an indication of low flow to the analyser or blockage in the sample line. Check for any blockages (such as water ingress), replace filters where necessary and check the flow indication is at a satisfactory level (approx. 200-300ml/min). If problems still persist please contact our technical support team for further assistance.
Why is the reading not as I expected? | Is auto-calibration operating correctly? Any system faults including low calibration pressure, can be viewed from the home screen. Is the auto-calibration frequency sufficient for the application?
What concentration of calibration gas do I need to connect to the system? | For grid injection applications, we recommend 100% CH4, however, this can be altered and configured locally if using the system to measure gas in various other applications.

## Slide 3: Biomethane 3000 FAQ’s
Question | Answer
What size cylinder is required? | This is entirely up to the user, the system can be configured to operate using 110 litre check gas cylinders or if using a larger external cylinder, a regulator would need to be provided with tube fittings ready to accommodate 6mm tubing.
How often do I need to replace the cylinder? | This is dependent on frequency of calibration, the system has the facility to notify the user when the cylinder is below a certain pressure, giving plenty of time to arrange a replacement form the local provider.
How often does the calibration routine run? | This is user definable, from as frequent as hourly to daily.
How can I check the calibration gas pressure? | The auto-calibration gas bottle pressure is available as a Modbus output (see manual). Any faults relating to bottle pressure can be viewed from the “View Faults” option on the Home screen of the analyser.